# Reference dataset

This reference dataset has been designed in order to be distributed along with the NCS report on Quality assurance of Treatment Planning Systems and its specific suggested tests concerning the patient and dose modeling. It has been created using MATLAB (R2017b, MathWorks Inc, Natick, MA, USA), especially the DICOM Manipulation Tools developed by Mark Geurts from University of Wisconsin [geurts], and respecting the following procedure for the different UIDs [newhauser].

The figures have been generated using 3D Slicer (v4.8.1, Surgical Planning Laboratory, Harvard Medical School, Boston, MA, USA) [fedorov], and especially its RT module, SlicerRT from Queen’s University (Ontario, Canada) [pinter].

## Patient information

The patient’s information has been designed as follows.

Family name		NCS TPS QA

Given name		PHANTOM

ID			NCSTPSQA2020

Birth date		18/04/2018

Sex			M

Institution name	NCS

## CT1

The first phantom is a CT of 280 slices of 512x512 voxels. The voxel size is 1x1x1 mm³. The origin is set at the center of the image. The ‘ImagePatientPosition’, which is the position of the center of the upper left voxel, is at (-255.5, -255.5, 139.5) mm for the first image.

Here is a summary of the relevant Dicom tags of this phantom.

Directory			phantom1

Files				CT1_xxx.dcm

Modality			CT

Series description		NCS TPS QA Phantom 1

Patient position			HFS

Voxel size [mm³]		1 x 1 x 1

Image size [voxels]		512 x 512 x 280

Image 1 Patient position [mm]	(-255.5, -255.5, 139.5)

Some structures have been set to a known Hounsfield unit as the background is set to -1000.

Position	Dimensions			HU	Mimicking tissue

Cylinder	Centered	H = 240 mm and R = 200 mm	0	Water

Big cube	Centered	L= 200 mm			-100	Fat

Cube 1		Centered	L = 20 mm			40	Soft tissue

Cube 2		50 mm LEFT	L = 20 mm			1800	Cortical bone

Cube 3		70 mm POST	L = 20 mm 			4000	Prosthesis

Cube 4		40 mm SUP	L = 20 mm			-700	Lung

## RTStruct

Directory			phantom1

File				RTSS1.dcm

Modality			RTSTRUCT

Series description		NCS TPS QA RTSS1

Structure set label		RTSS1

Manufacturer model name	WriteDICOMStructures

A RTStruct file has been generated for all the structures visible in CT1 on every slice. All structures have a geometric type set to ‘CLOSED_PLANAR’, a contour slab thickness of 1 mm and different Dicom interpreted types, as summarized in the following table.

#	Name		Type		Slices (Top, Bottom) [mm]	Points/slice	Volume	 [cm³]

1	ProsthesisCube	PTV		20 (9.5, -9.5)			40		8.0

2	FatCube	ORGAN		200 (99.5, -99.5)		400		8000.0

3	LungCube	GTV		20 (49.5, 30.5)			40		8.0

4	BoneCube	PTV		20 (9.5, -9.5)			40		8.0

5	SoftCube	AVOIDANCE	20 (9.5, -9.5)			40		8.0

6	WaterCylinder	EXTERNAL	240 (119.5, -119.5)		201		30159.3

## RTPlan and RTDose

Directory			phantom1		phantom1

Files				RTPlan1.dcm		RTDose1.dcm

Modality			RTPLAN			RTDOSE

Series description		NCS TPS QA RTPlan	NCS TPS QA RTDose

Manufacturer model name	WriteDICOMPlan	WriteDICOMDose

The RTPlan is composed of 2 opposed static photon beams with open fields of 40x40 cm² for 5 fractions of treatment. Some relevant DICOM tags are summarized in the following table.

Beam name		ant		post

Beam type		STATIC		STATIC

Radiation type		PHOTON	PHOTON

Delivery type		TREATMENT	TREATMENT

Gantry angle [°]		0		180

Collimator angle [°]	0		180

Beam meterset [MU]	1644.4		1770.7

Dose rate [MU/min]	400		400

The RTDose has been designed to match exactly the matrix of the primary image (CT1). The dose grid size is 1x1x1 mm³ and the ‘ImagePositionPatient’, which is the position of the upper left voxel of the last slice in RTDose files, is (-255.5, -255.5, -139.5) mm. The units and type of dose are set to ‘GY’ and ‘PHYSICAL’ respectively. In this file, the structures have been set a uniform dose as illustrated in the table below.

Structure	Dose [Gy]

BoneCube	25

ProsthesisCube	30

LungCube	35

SoftCube	40

FatCube	20

WaterCylinder	5

The dose volume histogram (DVH) does not seem to be perfectly displayed in this viewer. For the 4 little cubes - BoneCube, ProsthesisCube, LungCube and SoftCube - there should only be a perfect step at 25, 30, 35 and 40 Gy respectively, as represented in the following figure. As the FatCube and the WaterCylinder encompass other structures, their DVH should not be a perfect step-like curve. Some dose statistics are summarized in the following table.

Structure	Dmin [Gy]	Dmax [Gy]	Dmean [Gy]

BoneCube	25		25		25

ProsthesisCube	30		30		30

LungCube	35		35		35

SoftCube	40		40		40

FatCube	20		40		20.050

WaterCylinder	5		40		8.992

## MRI

Directory			phantom1MRT2

Files				MRT2_xxx.dcm

Modality			MR

Series description		NCS TPS QA Phantom MR T2

Patient position			HFS

Voxel size [mm³]		1 x 1 x 1

Image size [voxels]		512 x 512 x 280

Image 1 Patient position [mm]	(-255.5, -255.5, 139.5)

An MR T2-like image has been created, using the same resolution and position as the CT1, on a 3T imager. Signal values have been assigned to the different structures but they do not reflect any physical properties.

Structure	Signal value

SoftCube	100

LungCube	10

BoneCube	300

ProsthesisCube	0

FatCube	400

WaterCylinder	700

## PET

Directory			phantom1PET

Files				PET_xxx.dcm

Modality			PT

Series description		NCS TPS QA Phantom PET FDG

Study description		18F-FDG PET

Patient position			HFS

Units				BQML

Voxel size [mm³]		2.56 x 2.56 x 2.80

Image size [voxels]		200 x 200 x 100

Image 1 Patient position [mm]	(-254.72, -254.72, 138.60)

An FDG PET-like image has been created, using first the same resolution and size as CT1 (1x1x1 mm³ and 512x512x280 voxels). Activities have been assigned to the different structures, as shown in the following table, with the background set to 0.

Structure	Activity [Bq/ml]

SoftCube	24000

LungCube	8000

BoneCube	8000

ProsthesisCube	8000

FatCube	3000

WaterCylinder	500

The 3D matrix has then been resampled to 200x200x100 voxels, leading to a new voxel size of 2.56x2.56x2.80 mm³, but keeping the origin in the center of the phantom with a new ‘ImagePositionPatient’ set to (-254.72, -254.72, 138.60) mm.

## Patient position

New CTs have been generated based on CT1 to check for patient position and image orientation. The size and resolution were kept identical.

Directory		phantom1FFS			phantom1HFP		phantom1FFP

Files			CT1_FFS_xxx.dcm		CT1_HFP_xxx.dcm	CT1_FFP_xxx.dcm

Series Description	NCS TPS QA Phantom 1 FFS	… HFP			… FFP

Patient position 	FFS				HFP			FFP

Image orientation	[-1 0 0 ; 0 1 0]			[-1 0 0 ; 0 -1 0]		[1 0 0 ; 0 -1 0]

## Shifts and tilts

Directory			phantom1Shift

Files				CT1_shift_xxx.dcm

Modality			CT

Series Description		NCS TPS QA Phantom 1 Shifted

Voxel size [mm³]		1 x 1 x 1

Image size [voxels]		512 x 512 x 280

Image 1 Patient position [mm]	(-255.5, -255.5, 139.5)

A shifted CT has been created based on CT1. The size and the resolution are kept the same. Only the structures in the image were shifted according to the following table. The ‘ImagePositionPatient’ is kept at (-255.5, -255.5, 139.5) mm.

Shift [mm]	Direction

AP	20		POST

LR	30		LEFT

SI	10		INF

A rigid registration Dicom file has also been created, bringing back the CT shifted to the space of CT1.

Directory		phantom1Shift

File			REG_shift.dcm

Modality		REG

Series description	NCS TPS QA REG shifts

The transformation matrix is naturally given by

10010-300-2000001   100   1

which is a translation vector of (-30, -20, 10) mm and a rotation vector of (0, 0, 0) °.

Directory			phantom1Tilt			phantom1MRtilt

Files				CT1_tilt_xxx.dcm		MRtilt_xxx.dcm

Modality			CT				MR

Series description		NCS TPS QA Phantom 1 Tilted	NCS TPS QA Phantom MR Tilted

Voxel size [mm³]		1 x 1 x 1			2 x 2 x 4

Image size [voxels]		512 x 512 x 280			256 x 256 x 70

Image 1 Patient position [mm]	(-255.5, -349.0, 93.5)		(-255.0, -347.8, 92.5)

Image n Patient position [mm]	(-255.5, -209.5, -148.1)		(-255.0, -209.8, -146.6)

A tilted CT has also been created based on CT1, and a tilted MRI as well. The CT kept its original resolution and size while it changed to 256x256x70 voxels and 2x2x4 mm³ for the MRI, without changing the origin of the images.

Both images have been tilted with an angle of 30° along the LR axis. The ‘ImagePositionPatient’ should also vary, not only through its slice location, but taking tilt into account as well.

A rigid registration Dicom file has also been created, bringing back the CT tilted to the space of CT1.

Directory		phantom1Tilt

File			REG_tilt.dcm

Modality		REG

Series description	NCS TPS QA REG tilt

The transformation matrix is naturally given by

1000.87000.5000-0.50000.87001

which is a translation vector of (0, 0, 0) mm and a rotation vector of (-30, 0, 0) °.

## Dose grid

Other RTPlans and RTDoses were created to test for the dose grid effect. The RTPlans have the same configuration as the original RTPlan.

Directory		phantom1rt2			phantom1rt3

Files			RTPlan2_3mm.dcm		RTPlan3_1mm.dcm

Modality		RTPLAN				RTPLAN

Series description	NCS TPS QA RTPlan2 3mm	NCS TPS QA RTPlan3 1mm

The original RTDose has been designed to match exactly the matrix of the primary image (CT1). The dose grid size is 1x1x1 mm³ and the ‘ImagePositionPatient’, which is the position of the upper left voxel of the last slice in RTDose files, is (-255.5, -255.5, -139.5) mm. The units and type of dose are set to ‘GY’ and ‘PHYSICAL’ respectively. In this file, the doses in the structures have been set to produce the following DVH.

This original RTDose was first resized to a 170x170x92 voxels matrix with a dose grid of 3x3x3 mm³, leading to the RTDose2_3mm file. This dose was then super-sampled to a 510x510x276 voxels matrix with a dose grid of 1x1x1 mm³, the RTDose3_1mm file.

Directory			phantom1rt2			phantom1rt3

Files				RTDose2_3mm.dcm		RTDose3_1mm.dcm

Modality			RTDOSE			RTDOSE

Series description		NCS TPS QA RTDose2 3mm	NCS TPS QA RTDose3 1mm

Voxel size [mm³]		3 x 3 x 3			1 x 1 x 1

Image size [voxels]		170 x 170 x 92			510 x 510 x 276

Image Patient Position [mm]	(-253.5, -253.5, -136.5)		(-254.5, -254.5, -137.5)

This new dose distribution produces a DVH as illustrated in the following figure. Some dose statistics are summarized in the following table.

Doses (Gy)	Min	Max	Mean	D1%	D5%	D20%	D50%	D80%	D95%	D99%

SoftCube	14.89	50.40	31.29	16.02	18.85	20.40	32.88	36.64	43.31	50.40

LungCube	11.80	44.61	27.23	13.12	16.11	17.74	28.48	32.77	37.13	44.61

ProsthesisCube	10.04	38.48	23.33	11.30	13.42	16.34	24.15	28.63	32.79	38.48

BoneCube	11.17	32.28	20.03	12.25	13.10	16.21	19.96	23.43	27.42	32.28

FatCube	2.02	50.40	8.83	2.42	6.59	7.69	8.65	9.96	11.87	13.97

WaterCylinder	0.05	50.40	3.78	0.86	1.83	2.00	2.00	7.85	10.04	12.26

## CT2

Another phantom has been designed to test for structure expansion and shrinkage. This phantom is a CT of 300 slices of 512x512 voxels. The voxel size is 1x1x1 mm³. The origin is set at the center of the image. The ‘ImagePatientPosition’ is at (-255.5, -255.5, 149.5) mm for the first image.

Directory			phantom2

Files				CT2_xxx.dcm

Modality			CT

Series description		NCS TPS QA Phantom 2

Patient position			HFS

Voxel size [mm³]		1 x 1 x 1

Image size [voxels]		512 x 512 x 300

Image 1 Patient position [mm]	(-255.5, -255.5, 149.5)

The diabolo structures have been designed following the instructions of the NCS report 15 [ncs15] and are recalled here.

The different structures are centered and have been set to a known Hounsfield unit as the background is set to -1000.

Dimensions (mm)		Volume (cm³)	HU	Mimicking tissue

Cylinder	H = 280, R = 200		35175.6	0	Water

Diabolo1	R1(0) = 	40, R2 = 150, m = 0	1777.0		40	Soft tissue

Diabolo+	R1(0) = 	40, R2 = 150, m = 10	3188.1		-700	Lung

Diabolo-	R1(0) = 	40, R2 = 150, m = -5	1281.6		1800	Cortical bone

## References

[geurts] Geurts M. DICOM Manipulation Tools. 2018, GitHub Repository, University of Wisconsin, Madison, WI, USA, https://github.com/mwgeurts/dicom_tools

[newhauser] Newhauser W, Jones T, Swerdloff S, Newhauser W, Cilia M, Carver R, Halloran A, Zhang R. Anonymization of DICOM electronic medical records for radiation therapy. 2014, Computers in Biology and Medicine 53: 134-140

[fedorov] Fedorov A, Beichel R, Kalpathy-Cramer J, Finet J, Fillion-Robin J-C, Pujol S, Bauer C, Jennings D, Fennessy FM, Sonka M, Buatti J, Aylward SR, Miller JV, Pieper S, Kikinis R. 3D Slicer as an Image Computing Platform for the Quantitative Imaging Network. 2012, Magnetic Resonance Imaging 30(9):1323-1341

[pinter] Pinter C, Lasso A, Wang A, Jaffray D, Fichtinger G. SlicerRT – Radiation therapy research toolkit for 3D Slicer. 2012, Medical Physics 39(10):6332-6338

[ncs15] Bruinvis IAD, Keus RB, Lenglet WJM, Meijer GJ, Mijnheer BJ, van ‘t Veld AA, Venselaar JLM, Welleweerd J, Woudstra E. Quality assurance of 3-D treatment planning systems for external photon and electron beams. Report 15 of the Netherlands Commission on Radiation Dosimetry. 2005